## Slide 1
Vishwakarma Application ID
VK26-IITI-HT-0050
Dermasense: a laser-written graphene strip that tells nurses when a wound dressing is full
Soham Prabhakar Kadu · Bhakti Ajay Kadam | Health Tech: Emergency Response & Remote Monitoring

## Slide 2
Problem Context
Core Insight: dressing makers already define 'full' as a distance from the pad edge. Dermasense measures exactly that line, from outside the wound.
44.9 %
OF CHANGES MADE AT OPTIMAL MOISTURE¹
≈ 2 cm
MAKER'S 'CHANGE NOW' LINE FROM PAD EDGE²
1 spot
WHAT TODAY'S MOISTURE SENSOR READS¹
1 laser pass
TO WRITE A GRAPHENE ELECTRODE³
Figure: exudate spreads outward; makers say change when it nears 2 cm from the pad edge², but under a bandage nobody can see it
Target Beneficiary & Before-State User Flow
Target User: wound-care nurses (clinics, home care) and patients with exuding chronic wounds: diabetic foot, pressure and venous leg ulcers.
Operational Bottleneck: the 'change now' signal is visual. Under bandages or opaque covers it is hidden, so dressings are changed by the clock or after a leak.
Before-State Flow: dressing on → exudate spreads unseen → changed too early (wasted dressing, disturbed wound) or too late (leak, macerated skin) → repeat.
Public Search Disclosure: WoundSense reads moisture at one spot under the dressing¹; a 2007 lab array mapped moisture⁴; Stanford's smart bandage uses cleanroom electronics⁵; graphene wound sensors are still lab-only⁶. None reads the maker's change line at rupee cost.
▪
▪
▪
▪

## Slide 3
User Journey
Design Principle: the strip sits on the outer face of the absorbent pad, never on the wound: lower biocompatibility risk and no fouling of the sensor.
By clock
HOW DRESSINGS ARE CHANGED TODAY
4 zones
CENTRE → EDGE ON ONE STRIP
0
CONTACT WITH THE WOUND BED
≤ ₹100
TARGET COST PER STRIP
Figure: today's blind cycle vs the Dermasense cycle (target workflow, to be proven on the bench first)
After-State User Flow & Quantified Efficiency
After-State Flow: strip on the pad, reader clipped on → zones read wet/dry → phone shows OK, SOON, CHANGE or LEAK → change when the change line is wet.
Who Benefits: patients get fewer needless removals; nurses get a reason for every change; clinics waste fewer dressings.
Efficiency Target: alert within ±10 % of the fluid volume at which exudate visibly reaches the change line, in bench tests with simulated exudate⁹.
Not Claimed Yet: fewer dressing changes or faster healing. Those need clinical studies after the bench prototype works.
▪
▪
▪
▪

## Slide 4
System Architecture & Specs
Architecture Rationale: only film and graphene are thrown away; every electronic part lives in the reusable reader.
Figure: disposable LIG strip on the pad + reusable clip-on reader + phone (concept illustration; not yet built)
Parameter | Target Spec | Selected Hardware / Protocol
Sensing strip | 4 wet/dry zones | Laser-induced graphene on Kapton
Zone signal | Dry/wet ≥ 100× (target) | AC drive between LIG electrodes
Reader | All zones every 5 s | ESP32-C3 + ADS1115 (16-bit)
Temperature ref. | ±0.1 °C class | TMP117, I²C
Display | OK / SOON / CHANGE / LEAK | USB log now, BLE app next
Strip cost | ≤ ₹100 (target) | One laser pass, no masks
Engineering Target Specifications & Subsystems

## Slide 5
CAD Design & Mechanical Specs
Why Process First: every graphene sensor needs a repeatable laser recipe, and '12' sits in the low-power range where CO₂ tubes can be unstable⁷.
Figure: our laser test coupon (real DXF from make_test_pattern.py) and the E1 settings around 'setting 12'
Design Files & Laser Process Window
Test Coupon: 6 cells × 3 graphene strips, 2 mm wide, 8/16/24 mm long; each colour is one laser setting.
Measurement Trick: three lengths separate the graphene's sheet resistance from probe contact resistance, so a multimeter is enough.
Reference: 'setting 12', where we saw a graphene-like layer, appears twice per coupon, on 2 days.
Accept Rule: the recipe passes if sheet resistance varies ≤ 20 % between runs, fixed before testing.
Next in SolidWorks: zone strip, clip-on reader and dressing test rig, sized from E1's measured line width.
▪
▪
▪
▪
▪

## Slide 6
Electronics Schematics & BOM
Procurement Status: every part is available in India. Prices are estimates or listings seen on 14 Oct 2026; we order after E1 confirms the laser recipe.
Figure: zone readout; each zone is driven both ways and read by a 16-bit ADC (bench design; KiCad capture next)
Component | Part / Specification | Qty | Lead Time | Cost (INR)
Microcontroller | ESP32-C3 dev board | 1 | Local | ₹250–700
ADC | ADS1115 16-bit, I²C | 1 | Local | ₹200–450
Temp. reference | TMP117 breakout | 1 | Online | ₹446
Resistors | 100 kΩ 0.1 %, 1 kΩ | 4 | Local | ₹40–160
Pad contacts | Silver paint + Cu tape | 1+1 | Local | ₹750–1,800
Test fluid | NaCl + CaCl₂ (Solution A) | Lab | Lab stock | ₹0
Total (excl. film) | ₹1.7–3.6k |  |  | 
Itemized Prototype Bill of Materials (BOM)

## Slide 7
Maker Journey
Team: Soham Prabhakar Kadu & Bhakti Ajay Kadam. Faculty mentor sign-off to be added after the E1 review.
Figure: our real work so far: coupon design file, E1 data sheet, machine settings read from our LaserCAD install (lab photos follow E1)
Stage 1 to Stage 2 Engineering Iterations
Iterative Pivot: Stage 1's closed-loop electrotherapy bandage was too broad to build well; we narrowed it to one measurable job: telling when a dressing is full.
Laser Evidence: our earlier trial on Kapton at setting '12' gave a black, graphene-like layer; E1 repeats it with full records.
Machine Facts: LaserCAD V8.28 on our laptop: 1200 × 900 mm bed, 20 kHz PWM, 98 % cap. So '12' is a % command, not watts.
Repository: design scripts, coupon DXF, E1 protocol and data sheets under version control with timestamps (link to be added).
▪
▪
▪
▪

## Slide 8
Safety by Design & Roadmap
Failure Mode | Potential Hazard | Embedded Fail-Safe / Mitigation Response
Laser over-power | Burn-through, flame, polyimide fumes | Lowest setting first; exhaust on; never unattended; stop on any flame
Unstable low-power output | Graphene quality varies run to run | Log tube current; reference cell ×2 per coupon; repeat on 2 days
Electrodes near skin (future) | Irritation, current through tissue | Strip stays off the wound; AC only, ≤ 10 µA; ISO 10993 tests before skin use
Design FMEA (Failure Mode and Effects Analysis)
Fabrication Phase (Months 1 - 2)
Oct 2026: E1: reproducible graphene on our laser; measure sheet resistance and line width (4 coupons, 2 days).
Nov 2026: zone strip v0.1 cut from E1 dimensions; bench reader (ESP32-C3 + ADS1115) assembled and calibrated.
Integration & Bench Demo (Months 3 - 4)
Dec 2026: E3: strips on real dressings with simulated exudate⁹; zone alerts vs the visible stain front.
Jan 2027: repeatability and bending tests; Stage 3 demo video with raw data.
▪
▪
▪
▪

## Slide 9
Technical Challenges & Risks
Fallback: if our laser cannot give repeatable graphene, we use a partner lab's laser; the strip design and reader stay the same.
Observed Risk: '12' is a low-power command on a 100 W tube; such tubes fire unevenly below ~25 %⁷, and the polyimide brand changes the result⁸.
Mitigation: log tube current, repeat the reference on 2 days, accept only if variation ≤ 20 %; fall back to slower or defocused passes.
Test: E1, 8 reference cells over 2 days, tube current logged.
Observed Risk: black and conductive is not proof of graphene; resistance alone cannot show graphitic carbon.
Mitigation: Raman (D, G, 2D peaks) and SEM through a university facility; until then we call it a graphene-like layer.
Test: Raman on our 3 best cells; graphene shows D ≈ 1350, G ≈ 1580 and 2D ≈ 2700 cm⁻¹ bands³.
Wound Contact: none by design; the strip sits on the outer face of the pad.
ISO 10993: required before any skin contact; Kapton is not sold as medical grade¹⁰.
IEC 60601-1: applies to the future on-body reader, not to bench tests.
Claims: bench data only, so no diagnostic or healing claims.
Laser Repeatability
Is It Graphene?
Safety & Standards

## Slide 10
1. Milne S.D. et al., A wearable wound moisture sensor as an indicator for wound dressing change, Int. Wound J. 13(6):1309–1314 (2016). doi:10.1111/iwj.12521
2. Dressing instructions for use: Mölnlycke Mepilex Border Flex 'time to change'; 3M Tegaderm silicone foam border guide; foam dressings citing the 2 cm rule.
3. Lin J. et al., Laser-induced porous graphene films from commercial polymers, Nat. Commun. 5, 5714 (2014). doi:10.1038/ncomms6714
4. McColl D., Cartlidge B., Connolly P., Real-time monitoring of moisture levels in wound dressings in vitro, Int. J. Surg. 5(5):316–322 (2007). doi:10.1016/j.ijsu.2007.02.008
5. Jiang Y. et al., Wireless, closed-loop, smart bandage with integrated sensors and stimulators, Nat. Biotechnol. (2022). doi:10.1038/s41587-022-01528-3
6. Asgharian H. et al., WoundMx: wound-infection biomarker detection with laser-induced graphene, Research Square preprint (2025). doi:10.21203/rs.3.rs-7274589/v1
7. CO₂ laser tube current vs PWM setting, engineering measurements, softsolder.com (2022), not peer-reviewed.
8. Influence of lasing parameters on the morphology and electrical resistance of polyimide-based LIG, Carbon Letters (2023).
9. EN 13726-1:2002 / EN 13726:2023, Test methods for primary wound dressings (absorbency; 'Solution A' simulated exudate).
10. DuPont statement on the use of DuPont materials in medical applications (Kapton medical caution).
11. Reviewed, not used as evidence: Rezaei I. et al., Wearable Kapton graphene biosensor for toxic gases, J. Hazard. Mater. Adv. 15, 100452 (2024), simulation only.
References & Sources
(numbers match the superscripts on each slide)

## Slide 11
Image Sources
All figures were drawn by the team with Python from project files: problem, journey and concept illustrations; the coupon from lig_test_coupon.dxf; the schematic with schemdraw; machine settings from LaserCAD SysCfg.ini. Script: 08_submission/make_figures.py.
No stock photos and no AI-generated images are used; concept drawings are labelled as concepts.